Low-Resistant Stabilized Mini-Vaporizers IN (VIC) and OUT (VOC) of the Breathing Circuits (manuscript 2010)

I.V. Molchanov, V.A. Sidorov, L.L. Nikolaev and A.Z. Berlin

Russian Academy of Anesthesiology and Intensive Care

Department of Anesthesiology and Intensive Care (Filatov Children’s Hospital, Moscow, Russia)

Ural Optical&Mechanical Plant (Chief R&D medicine department I.K. Sergeev, Moscow, Russia)

MITC-M Ltd. (Manufacture Company - Director I.K. Gorlin, Moscow, Russia)

MINIVAP Ltd ( Scientific & Manufacture Company -  - Director A. Z. Berlin, Moscow, Russia)

# Abstract

The main problems related to inhalation anesthesia may be solved by using a more advanced vaporizer that would be accurate as a plenum vaporizer, simple and low-resistant as a draw-over one.

# Due to the low resistance and virtual independence from fresh gas flow rate, temperature, and ambient pressure, MINIVAP vaporizers are instantly adaptable to needs of both on-site emergency surgery and more sophisticated demands of a general hospital.

MINIVAP vaporizers have been successfully tested in two leading Moscow hospitals (> 100 anesthesias in adults and children 2-14 years) with different anesthesia machines (Drager Primus, STEPHAN Artec, Datex-Ohmeda S/5 Avance, Ru PMT Xena-010).

A MINIVAP vaporizer may be installed either IN a conventional breathing circuit (instant change of inspired concentration) or OUT of breathing circle (high safety) of any anesthesia machine.

Consumption of a liquid anesthetic by MINIVAP is approximately 5-10 ml/hour at mini or low-flow anesthesia (FG = 0.5-1 L/min) and 1-3 ml in the closed system (during 1- to 5-hour anesthesia).

Keywords

Anesthesia, low-flow, mini-flow; volatile anesthetic, inspired and expired concentration; fresh gas flow, vapor anesthetic; equipment, low-resistant (draw-over) vaporizers, plenum vaporizers; anesthesia machine, semi-closed circuit, closed circuit.

Introduction

Optimal inhalation anesthesia techniques include:

- Sevoflurane, isoflurane, or desflurane low- or mini-flow anesthesia [1, 2]

- Xenon combined anesthesia [3, 4].

Inhalation anesthesia is especially indicated for children up to 14 years old, who refuse any painful manipulations [5].

Analysis

Main problems of inhalation anesthesia include:

1 – The need for sophisticated and bulky anesthetic equipment (especially compared to what is used for intravenous anesthesia) [6].

2 - Non-stability of the well-known draw-over vaporizers (OMV, Ohmeda PAC, Goldman vaporizers are associated with unpredictable output and inadaptability to gas flows below 4 L/min) [7, 8].

3 – Delayed change of anesthetic concentration in closed and semi-closed systems, while the concentration inspired by the patient is not known with certainty in spite of an accurate VOC.

At the minimum oxygen flow 250 ml.min-1 the supply of “anesthetic agent reaches its limits…with isoflurane – and that would be more striking with enflurane or halothane – the individual uptake of an adult patient exceeds the amount of agent which maximally can be supplied into the breathing (VOC) system” [1].

The delivered anesthetic vapor flow may be increased by serially connecting agent-specific vaporizers kept out of the circuit [9]. On the other hand, the super-complex anesthesia machine (station) PhysioFlexTM with an electronically controlled vaporizer changes sevoflurane concentration from 0 to 2 vol.% in 80 to 510 s [10]. However, the mentioned methods and equipment are impractical.

There is a marked difference between the inspired anesthetic concentration and the corresponding vaporizer setting in circulation systems (Table 1). The lower the fresh gas flow, the higher the concentration to be dialed at the VOC to maintain a desired anesthesia depth.

Table 1

Data of inspired mixture of Sevoflurane (S) or Isoflurane (I) and O2 during mini- and low-flow anesthesia with a vaporizer OUT of the breathing circle [1]

| Fresh gas flow FG, L/min | 0.25 | 0.5 | 1 |
| --- | --- | --- | --- |
| Vaporizer setting / Inspired concentration, CS/ CI | S | 8/2.8 | 5/2.8 | 3/1.8 |
|  | I | 5/1.2 | 5/1.8 | 2/1 |
| Delivered anesthetic flow, FAD= CS (1-CS)-1 FG, mL/min | S | 21 | 26 (max 42) | 31 (max 84) |
|  | I | 13 | 26 | 20 (max 53) |

On the other side, the mentioned problems may be solved by using a perfect vaporizer that would be accurate as a plenum vaporizer, simple and low-resistant as a draw-over one. During induction of anesthesia (about 3 minutes) such a vaporizer installed IN the breathing circuit would enrich the entire inspired mixture at any fresh gas flow. For example, at a minute ventilation (MV) of 5 L/min and a vaporizer setting of 4 vol % the delivered anesthetic flow would be 200 mL/min instead of 13-42 mL/min of any VOC at the mini fresh gas flow (see Table 1).

An attempt to regularize the governing processes (gas flow, mass and heat transfer) within the laminar flow regime enabled us to design new low-resistant vaporizers (pocket MINIVAP-20 and stationary MINIVAP-100, Fig. 1, 2), which are capable of dosing out anesthetics at a flow range of 0.2-15 L·min-1 just like plenum vaporizers [11, 12].

Technical data of the new vaporizers and analogous devices are shown in Table 2.

Table 2

Novel Low-Resistance Vaporizers and its Plenum and Draw-Over Analogs

| Parameters | Vapor 2000 Drager Germany | Delta OMV Penlon UK | МВ-100 МВ-20 “MINIVAP” Ltd. Russia |
| --- | --- | --- | --- |
| Gas flow range, L/min | 0,25 – 15 | 0,2–15 3-15 | 0,2-15 0,2-10 |
| Temperature range, °C | 10 – 40 | 15– 35 18-22 | 5 - 35 |
| Atmospheric pressure, kPa | 100 ( 5 | 100 ( 5 | 70 - 110 |
| Anesthetic volume, mL | 360 | 250 50 | 150 40 |
| Wick volume (waste), mL | 60 | 60 10 | 5 3 |
| Pressure drop at 10 L/min, mm H2O | 1100 | 1000 10 | 20 10 |
| Angle of tilt, degrees | 30 | 10 30 | 90 180 |
| Weight, kg | 6,5 – 8,5 | 5,7 2 | 1,5 0,5 |

MINIVAP vaporizers conform to relevant Russian and international standards including F1850-00 “Anesthesia Workstation and Their Components” and F1054-01 “Conical Fittings”.

Results

We compared the delay, i.e. time (T) needed to change inspired anesthetic concentration using MINIVAP-20 vaporizer at OUT and IN positions of the breathing circuit of Drager (Primus), STEPHAN (Artec), Datex-Ohmeda (S/5 Avance), and Ru PMT (Xena-010) anesthesia machines (we used a bag or bellow instead of a patient). According to Conway`s formula [1]

T = VC . FG-1											(1)

Example 1. VOC. The concentration delivered by the vaporizer is diluted by the gas already contained in the circuit. At the best, (without a stagnant area) the delay (T) was > 10 min (at VC ≈ 5 L and FG = 0.5 L/min). After this time inspired, expired and delivered vaporizer concentrations will be approximately equal (within the tolerances).

VIC. The MINIVAP-20 vaporizer is installed between the inspired fitting 22M and the corrugated tube (volume VT ≈ 0.4 L) of any anesthesia machine (Fig. 2). The delay was much less (Fig. 3)

T = VT . MV-1 < 5 s										(2)

But in this case the anesthetic inspired concentration CI will be much higher than the scale one CS at a high minute ventilation and fresh gas flow [13, 5]

CI = CS (1-CS)-1 MV . FG-1  ≈ CS MV . FG-1							(3)

Anesthetic gas analyzers or Riken indicator with a water vapor absorber were used in all cases of VIC circulation systems.

MINIVAP vaporizers were successfully tested in two leading Moscow hospitals (more than 100 anesthesias delivered to adults and children aged 2 to 14 years, Tables 3, 4).

Routine anesthetic equipment and non-invasive monitoring were applied; readings were taken every minute during the surgical operation and every 2.5 or 5 min during maintenance of anesthesia.

Inspired and expired concentrations of CO2, O2, and N2O, as well as sevoflurane, isoflurane, or halothane, and respiratory parameters (inhaled and exhaled gas volumes, minute ventilation, inspired and expired pressure) were monitored.

Non-invasive monitoring included electrocardiography with pulse monitoring, arterial blood pressure measurement, and pulseoxymetry.

Table 3

Patients` (adults) and anesthesia characteristics (Botkin Municipal Hospital)

| Parameters | Urology | Surgery | Emergency surgery |
| --- | --- | --- | --- |
| Breathing circuit | Semi-closed (0.5-1 L/min), VIC | Semi-closed (1 L/min), VOC | Semi-closed, VOC and Semi-open |
| Anesthetic and duration of anesthesia (min) | Sevoflurane,90-150 | Sevoflurane, 90-150 | Isoflurane, 30-60 |
| Blood loss, ml | 100-2000 | 700-2000 | 100 |
| Age | 44-76 | 57-68 | 19-45 |
| Number of patients | 45 | 2 | 5 |

Table 4

Patients` (children) and anesthesia characteristics (Filatov Children’s Hospital)

| Parameters | Cystoscopy | Injuries (low-invasive) | Adenotomy and/or Tonsillectomy |
| --- | --- | --- | --- |
| Breathing components | Facemask | Laryngeal mask airway | Microlaryngeal tube |
| Breathing circuit | Semi-open (6 L/min) | Semi-open (6 L/min) | Semi-closed (0.5-1 L/min), VIC |
| Breathing | Spontaneous | Spontaneous | Artificial lung ventilation |
| Anesthetic and duration of anesthesia (min) | Sevoflurane or isoflurane, 15-20 | Isoflurane or halothane, 30-60 | Halothane or sevoflurane, 30-50 |
| Age | 9-14 | 10-14 | 2-8 |
| Number of patients | 13 | 15 | 64 |

There is depth of anesthesia III1 to control the administration of anesthetic drugs.

Example 2. A 5-year-old child (body weight 20 kg) had adenotomy under endotracheal anesthesia. After induction (N2O/O2 = 2:1 + halothane CS = CI = 3 vol %) with a high FG = 6 L/min during 2 min, tracheal intubation was performed and artificial lung ventilation was carried out throughout the anesthesia (tidal volume= 160 ml, rate = 17 cycles/min, MV = 0.16x17 = 2.7 L/min). Maintenance: N2O/O2 = 2:1 + halothane CS  = 0.3 % with FG = 1 L/min and MINIVAP-20 vaporizer IN the semi-closed circuit of Primus anesthesia machine. In that case the inspired concentration, as follows from equation (3), was CI ≈ 0.3 . 2.7 / 1 ≈  0.8 vol %.

In case of the minimal FG = 0.5 L/min, the vaporizer scale mark would be set as

CS≈ 0.8 . 0.5 / 2.7 ≈ 0.15 vol %.

Monitored inspired anesthetic concentration matched the calculated one with a total tolerance <0.45 vol.%, n = 34 (equation 3).

No complications or adverse reactions were observed (n = 114).

To deepen anesthesia quickly without increasing fresh gas flow, the vaporizer handwheel should be set to the maximum and the delivered anesthetic flow will be

FAD = CS . MV                                                                                                                                (4)

Example 3. According to the data from Example 2 and CS = 3 %, maximum VIC delivered anesthetic flow FAD = 0.03 x 2700 = 81 sm3/min during 0.5 min. After that the handwheel was set back to CS = 0,3 % with CI ≈  0.8 vol %.

The delay (see Fig. 3) T = VT / MV ≈ 0.4 / 2.7 ≈ 0.15 min < 10 sec.

To deliver to a patient the same anesthetic vapor volume at FG =0.5 L/min, the VOC should at first fill up the circuit volume V = 5 L during T1 = 5 / 0.5 = 10 min) and after that work during T2 = 2.7 min.

The total work time of the VOC T∑ ≈ 13 min compared with 0.5 min of the VIC.

The actual delay of VOC systems is much longer because anesthetic uptake by the patient.

To improve the efficiency of xenon anesthesia in surgical practice (allowed by the resolution of the Russian Health Ministry No. 363, 1999), a total of 22 laboratory tests with xenon and isoflurane combined anesthesia have been performed.

Example 4. Anesthesia was delivered to a dog weighing 25 kg. A “MINIVAP-20” vaporizer was installed OUT of the closed circuit of Artec (STEPHAN) anesthesia machine. After stabilizing the Xe:O2 ratio (70:30) the Xe flow was increased to 50-70 ml/min  and the O2  flow was increased 10% above the initial level. When Xe inspired concentration decreased to 50-60 vol.% owing to the N2 tissue elimination, we delivered isoflurane (4 times 2 minutes each during a 2-hour anesthesia with the fresh gas flow=0.25 L.min-1 and the vaporizer set to 3 %). The inspired isoflurane concentration was stabilized at 0.3-0.5 vol %, and expired one at 0.25-0.4 vol %.

The calculated consumption of liquid anesthetic by MINIVAP-20 during 2-hour anesthesia

# FAD = 5.1 FG T CS (1- CS )-1 = 0.64 ml = 0.96 g	(5)

corresponds to the vaporizer mass difference of 1.2 g measured before and after the anesthesia (differential tolerance = 0.3 g).

There were the same results when the “MINIVAP-20” vaporizer was installed IN the closed circuit of Russian PMT (Xena-010) anesthesia machine.

# Consumption of liquid anesthetic by MINIVAP vaporizer corresponds to the calculation equation (5) [13, 14]: at high fresh gas flow (6 L/min) it is about 20 ml/hour, while during mini- or low-flow anesthesia (0.5-1 L/min) it does not exceed 5-10 ml/hour, and in a closed system it is about 1-3 ml during 1-5 hours.

There is a minimum anesthetic wick waste: 3 ml of the MINIVAP-20 compared with 60 ml of the popular analogs (Table 2).

Anesthesia machine modifications

# Due to the low resistance and virtual independence from fresh gas flow rate, temperature, and ambient pressure, MINIVAP vaporizers are instantly adaptable to needs of both on-site emergency surgery and more sophisticated demands of a general hospital.

Field conditions

The MINIVAP Anesthesia Kit may be used in military surgery, urgent situations and remote areas, at district hospitals, and in veterinary anesthesia:

- at spontaneous breath and artificial respiration with air or oxygen (Fig. 6);

- with a low pressure oxygen source (concentrator);

- with “to-and-fro” and circle absorption systems;

- during Auto Analgesia [15].

The simplest modification (see Fig. 6) is the pocket vaporizer MINIVAP-20 (mass < 0.5 kg) that can be connected directly to a facemask or a tracheal tube by means of a non-rebreathing valve.

Hospital conditions

The vaporizer may be used OUT of the breathing circuit (VOC) instead of any other vaporizer at convenient fresh gas flow from 0.2 to 10 l/min (Fig. 8).

MINIVAP-20 may be installed IN the breathing circle of any anesthesia machine (Fig. 7).

It is possible to instantly switch over between In and Out positions of MINIVAP-20 with a special controller (Fig. 9), [16, 17].

Conclusions

# 1. Due to the low resistance and virtual independence from fresh gas flow rate, temperature, and ambient pressure, MINIVAP vaporizers are instantly adaptable to needs of both on-site emergency surgery and more sophisticated demands of a general hospital.

2. MV-20 vaporizer may be installed either IN a conventional breathing circuit (instant change of inspired concentration) or OUT of breathing circle (high safety) of any anesthesia machine.

3. Delay in changing the anesthetic inspired concentration VIC of closed and semi-closed systems 10-100 times smaller than the VOC.

References

1 Baum JA. Low Flow Anesthesia. Drager 2004.

2. Rupp K. Pediatric Anesthesia. 2005. Drager Medical AG & Co. KG.

3. Burov N.E., Molchanov I.V., Potapov V.N. and Nikolaev L.L. Trends in Development of Equipment for Xenon Anesthesia. Biomedical Engineering. 2005, 39, No. 6, 270-274.

4. Sidorov V.A., Tsipin L.E., Grebennikov V.A. Pediatric Inhalation Anesthesia. Moscow: MIA. 2010 – 184 p.p.

5. Mikhel`son V.A., Agavelyan E.G., Berlin A.Z., Sidorov V.A., Leshkevich A.I. Pediatric Inhalation Anesthesia: Stabilized Low Resistance Vaporizers In and Out of the Circuit.

Pediatric Surgery (Moscow). 2008; No 4: 51-56.

www.rusanesth.com

6. Davey A, Moyle JTB and Ward CS. Ward’s Anaesthetic Equipment. 3rd ed. 1992.

7. Watney G. In- and Out of Circuit Vaporizers. Anesthesia Equipment Resources ASE 2007. www.asevet.com/resources/circuits/circle.htm

8. Dobson MB. Anaesthesia at the District Hospital. 2nd ed. 2001.

9. Otero PE, Rebuelto M, Hallu RE, Aldrete JA. Increasing Anesthetic Concentration without Modifying FGF by Serial Connection of Vaporizers. ALFA 2003; 1-7.

10. Suzuki A., Bito H., Sanjo Y et al. Evaluation of the PhysioFlexTM Closed-circuit Anaesthesia Machine. Eur J Anaesthesiol, 2000. 17, 6, 359-363.

11. Berlin AZ. Anesthetic Delivering Method and Vaporizer. Ru Patent No 2329832 (2008).

12. Berlin A.Z. Anesthetic Vaporizer. Ru Patent No 2329069 (2008).

13. Molchanov I.V., Berlin A.Z., Burov N.E., Gribachev S.V., Korolev A.I. Low Resistance Stabilized Mini-Vaporizers In- and Out of Breathing Circles. Clinical Anesthesiology and Resuscitation (Moscow). 2007, 4, No 5, 66-71.

14. Philip JH. Low Fresh Gas Flow Oxygen and Agent Considerations. Brigham and Women’s Hospital 1985-2004. p.p. 122-125.

15. White D, Royston B. Respiratory Feedback Effects on Vaporizers in Circle Systems. Anaesthesia, 1998; 53: 555-9.

16. Berlin A.Z., Burov N.E., Molchanov I.V., Sidorov V.A., Jushkin A.V. Anesthesia Method. Ru Patent Application No 2008100074 (2008).

17. Berlin A.Z., Gorlin I.K., Grinvald A.V., Nikolaev L.L., Nikolin A.S., Ferdman M.I. Anesthesia Machine. Ru Patent No 2372947 (2009).

Illustrations

Fig. 1 – Vaporizers “MINIVAP-20”.

Fig. 2 – Vaporizers “MINIVAP-100”.

Fig. 3 – Time of inspire anesthetic concentration change (delay) from 0,5 to 1 vol % in semi-closed circles after alteration of the vaporizer setting. (“patient” – bag or bellow)

CVIC – inspired concentration with VIC (MINIVAP-20, MV = 5 L/min, VT = 0.4 L);

CVOC 2, CVOC 1, CVOC 0,5 - inspired concentrations with VOC (VC = 5 L, FG = 2; 1 and 0,5 L/min accordingly).

Fig. 4 – Delivered concentration of the “MINIVAP-20” vaporizer in flow range.

Fig. 5 – Inspired sevofluran concentration of MINIVAP-20 vaporizer in the open circuit: the vaporizer inlet is open, the outlet is connected with a self-inflating bag (C1 = V + B); the vaporizer inlet is connected with a self-inflating bag, the outlet is connected directly to a non-rebreathing valve (C2 = B + V, see Fig. 6).

Fig. 6 - “MINIVAP” Anesthesia Kit (open circuit at spontaneous breathing and with self-inflating bag).

Fig. 7 – 8. “MINIVAP-20” IN and OUT of the circuit of  an Anesthesia Machine.

Fig. 9 - Instant switching over In to Out position of “MINIVAP-20” with the special Valve.

| Fig. 1 (М 1:1) | Fig. 2 (М 1:3) |
| --- | --- |
| Fig. 6 | Fig. 7 |
| Fig. 8 | Fig. 9 |

Fig. 3

Fig. 4

Fig. 5